## Slide 1: Design of Time zero(T0) counter
contents
  Characteristic of p+p collision
  Time of Flight method
  Systematic scintillator performance
  Scintillator selection
  PMT selection
  Simulation of T0 counter
  Schematic design of T0 counter
  Background by T0 counter
Ver 6.2J
Univ. of Tsukuba
Hiroshi Tsuruoka , Masaya Ono

## Slide 2: What’s different in p+p collision ?
p+p １００GeV(JAM)
Au+Au 100AGeV(JAM)
Charged dn/dy distribution
dnc/dy
Low multiplicity(≈1/200 of Au+Au collision).
rapidity
p+p １００GeV(FRITIOF)

## Slide 3: Beam-Beam Counter(BB) in p+p collisions.
Mean=1.2〔particles/event〕
No.of hits | % | σBBC[ps]
0 | 41 | ―
1 | 23 | 87
2 | 17 | 60
3 | 10 | 50
4 | 5 | 43
5 | 3 | 39
6 | 1 | 35
Hit Multiplicity of Beam Beam Counter
No.of hits | %
0 | 84
1 | 14
2 | 2
Hit Multiplicity of Time-of-Flight detector region (67°<θ<113°,0°<φ<45°）
Mean=0.18〔particles/event〕
Few charged particles hit Time of Flight.
Need “Trigger” for efficient measurement.
～40％
No Hit at BBC
～84％
No Hit at ToF
BBC can’t be used as start counter.

## Slide 4: Purpose of T0
Time zero(T0) counter
    Covered Time-of-Flight detector region.
    Time resolution <50ps
Purpose
  Provides trigger for hadron measurement and start timing for TOF measurements
BBC can’t be used as start counter.
Few charged particles hit Time of Flight.
Need “Trigger” for efficient measurement.
TRIGER=CLOCK×T0(×TOF)
preferable

## Slide 5: Time of Flight (ToF) method
Time of Flight method resolution Required 100ps for 4σ
Stop counter (Time-of-Flight detector) resolution. <80ps
Start counter resolution
For Au+Au collision
start counter= Beam Beam Counter ΔTstop=60ps achieved
Time of Flight (ToF) resolution
β：particle velocity determined by Time of Flight
ｐ：particle momentum determined by tracks in the magnetic field
Particles are identified by mass
∴ΔTstart required 50~60 ps

## Slide 6: How to measure the start , stop Time
PMT2
PMT1
Charged particle
x
L-x
t0
Measured time
Obtain hit time t0 , hit position X
Resolution t0 , x
Resolution t1 , t2
N: Number of photo- electron
Time resolution has been found to be dominated by Number of photo-electron.

## Slide 7: Systematic scintillator performance
(ref M.Kurata et al. / NIM A 349(1994)447-45)
1.Light yield (proportional photo-electron)
Light yield is decrease exponentially with the distance from PMT
λLA is proportional scintillator’s cross section
λLA:light attenuation length
2.PMT’s time resolution
λTD:time degradation length
ΔT degrades exponentially with the distance from the PMT
Large cross section scintillater has small ΔT.
BC404 scintillator

## Slide 8: Scintillator selection
BC404 plastic scintillator
Physical constant | value
Light output(%anthracene) | 68
Wavelength of maximum emission | 408nm
Decay constant of main component | 1.8ns
Bulk light attenuation length | 160cm
Refractive index | 1.58
Radiation length | 42.5cm
We calculate 4 size(1.5,2.0,2.5,3.0cm)scintillator’s ΔT, which refer to Kurata’s paper.
Result is next page.
Required performanse
1.for good ΔT
large light yield ⇒thicker scintillator
2.a little background effect for backyard detectors
a little conversion probability⇒ thiner scintillator
Where do we compromise?

## Slide 9: φcoverage of T0 counter
Pt(MeV/c) | Max degree | Min degree | Required φcoverage(Δφ) 1set coverage of T0 is 10°
200 | 248° | 118° | 130°
300 | 235° | 136° | 100°
400 | 230° | 145° | 85°
*Max degree・・・The maximum degree that plus charged particles can hit Time-of-Flight detector .(particles also through drift ,pad chamber)
*Min degree・・・The minimum degree that minus charged particles can hit Time-of-Flight detector .(particles also through drift ,pad chamber)
Time-of-Flight detector covers φ=168°～213°
e+:Pt=400MeV
e+:Pt=300MeV
e+:Pt=200MeV
e-:Pt=400MeV
e-:Pt=300MeV
e-:Pt=200MeV
T0 counter
Time-of-Flight detector
Ｘ
y
φ

## Slide 10: Background by T0 (conversion)
1.Radiation length
Scintillator(BC404):λ=43cm
Cross Section | Conversion probability
2㎝×2㎝ | 2.4％
2.5㎝×2.5㎝ | 3.0％
3.0㎝×3.0㎝ | 3.5％
GEANT simulation
Cross Section | Conversion Probability
2㎝×2㎝ | 2.5％
2.5cm×2.5cm | 2.8％
3.0cm×3.0cm | 3.4％
Generate γray which have (π0⇒2γ） momentum
Thickness | Conversion Probability(GEANT) | From Radiation length(λrad=43cm)
2.0cm | 2.5％ | 2.4％
2.5cm | 2.8％ | 3.0％
3.0cm | 3.4％ | 3.5％
Charged Multiplicity =primary +secondary
=0.18+0.004(JAM)
=0.26+0.006(FRITIOF)

## Slide 11: ここから下は要らないトラペ

## Slide 12: Scintillator selection(2)
Every size achieve 50ps at diamond cut scintillator
2,2.5,3.0cm achieve 50ps at optically polished scintillator
We will design a T0 prototype with 2.0cm diamond cut scintillator